R&D FORMULATION SHEET
Glucose Sensor Working Formulation

- Purpose

To define the standardized preparation procedure for stock solutions, enzyme aliquots, and final glucose sensor working formulation for R&D and formulation development.

- Required Chemicals and Materials

- PVP K-15 (Polyvinylpyrrolidone, MW 10,000)

- Phosphate Buffered Saline (PBS, pH 7.2)

- Polyethylene Glycol (PEG 20,000)

- Triton X-100

- Silicone Antifoaming

- HPLC Water

- Falcon tubes 15ml or 50 ml

- Pipettes and pipette tips

- Vortex machine

- Weighing machine

- Spatula

- Tissue paper

- Aluminium foil

- Tube stand

- Precautions

Perform all work in clean and dust-free conditions.

Maintain exact reagent concentrations; do not alter specified quantities.

Protect potassium ferricyanide from light and humidity.

Use formulation within 60 minutes after ferricyanide addition.

Keep enzyme protected from light and airtight during thawing.

Do not vortex or harshly mix after enzyme addition.

Close all stock bottles immediately after use.

Discard and restart if procedural deviation occurs.

Note: In case while taking HPLC water with pipette ensure that no air bubbles in that tip (in case it occurs air bubbles make sure throw that old tip and use new tip again).

- Storage Conditions

Stock solutions: Store at 4°C.

Enzyme aliquots: Store at −20°C.

Final formulation: Use immediately; maximum hold time 60 min protected from light.

- Enzyme Aliquot Preparation

Remove bulk enzyme vial from −20°C and keep at 4°C for 20 min.

Bring to room temperature for 20 min.

Make 20-25 aliquots of 25 mg into amber micro-centrifuge tubes [MCT].

Label aliquots and store at −20°C.

Record thaw count on bulk vial.

- Weighing of Powder Chemicals

- Switch on the weighing balance.

- Tear a fresh foil and wrap it to make a well.

- Put the foil and tear to make zero.

- Clean the spatula with tissue before using and after using.

- Chemicals PVP K-15, PEG 20K and PBS pH. 7.2 in a weighing machine.

- Add the weighed powders into the respective Falcon tubes.

- Add the required quantity (ml) of HPLC water into all Falcon tubes.

(Note : Don’t keep pipette directly into the HPLC bottle, take it into from a new falcon tube).

- Stock Preparation

Prepare the following stock solutions and store at 4°C after labeling.

| Stock Name | HPLC Water | Chemical Quantity |
| --- | --- | --- |
| PVP K15 | 10 mL | 0.5000 g |
| PEG 20K | 10 mL | 0.5000 g |
| PBS 2X | 10 mL | 0.01716 g |

- Add the weighed powders into the respective Falcon tubes.

- Mix well using a vortex machine.

- Working Formulation Preparation Precautions

Bring enzyme aliquot to 4°C for 20 min, then RT for 15–20 min.

Bring stock solutions to RT and vortex properly before use.

In fresh falcon tube, add formulation components as per table below.

Vortex base formulation 5–10 min before ferricyanide addition.

Add ferricyanide slowly and vortex until dissolved.

Transfer ferricyanide mixture into enzyme vial.

Mix gently for 3–4 min until dissolved

- Note: (Do Not vortex).

- Final Working Formulation

| Stock Name | Chemical Quantity |
| --- | --- |
| PBS 2X | 500 ul |
| PEG 20K | 500 ul |
| PVP K15 | 500 ul |
| HPLC | 3467 ul |
| Tx-100 | 25 ul |
| Anti-forming | 8 ul |

- After adding all five chemicals, vortex properly to ensure uniform mixing. It will be slightly milky.

- Ensure no leakage from Falcon tube caps.

- Tighten all chemical bottles immediately after use.

- Cover the tubes with aluminum foil.

- Store the prepared formulation in the refrigerator at 4°C.

- Reagent Preparation

| Chemical Name | Quantity |
| --- | --- |
| Ferricyandie | 85 / 87.5 mg |
| Glucose oxidase | 25 mg |

- Take out all stock from 4°C and keep at Room Temperature for 15–20 min to bring to Room Temperature.

- Take out a Glucose Oxidase vial from -20°C and keep at Room Temperature for 15–20 minutes. (Note: Do Not mix the enzyme if it is cold)

- Take out a fresh falcon tube (15 ml) & write “formulation & date”.

- Vortex Final Working Formulation for 5–10 sec.

- Weigh 85 mg of potassium Ferrocyanide using a micro weigh balance and transfer it into the amber MCT.

- Add the Ferrocyanide into final working formulation.

- Close the lid of the MCT and mix it well using a vortex Machine.

- Using a 1000 ul micropipette transfer the Ferrocyanide dissolved solution and transfer it into a 25 mg enzyme vial kept at RT.

- Mix it gently for 3-4 minutes. (Do not vortex)

- The formulation is ready to be used